# Operational Qualification Template

  - Objective

The objective of Operational Qualification (OQ) testing is to demonstrate that the solution is operable under normal expected conditions with acceptable and expected output results

  - Scope

OQ testing involves using a Nymi solution to perform day-to-day activities that are defined in the Nymi Design/Configuration Specifications document

  - Assumptions

    - A reliable network connection exists between each component in the environment

    - The tester has access to Nymi Connected Worker Platform documentation

  - Exclusions

n/a

- Acronyms and References

  - Acronyms and Definitions

| Acronyms and Definitions |  |
| --- | --- |
| AC | Acceptance Criteria |
| CFR | Code of Federal (US) Regulations |
| cGxP | Abbreviation which includes current Good Manufacturing, Clinical and Laboratory Practices |
| CS | Configuration Specification |
| CSV | Computer Systems Validation |
| DS | Design Specification |
| FS | Functional Specification |
| GUI | Graphical User Interface |
| IQ | Installation Qualification |
| LAN | Local Area Network |
| OQ | Operational Qualification |
| PQ | Performance Qualification |
| SOP | Standard Operating Procedure |
| URS | User Requirements Specification |

  - References

| References |  |
| --- | --- |
| 21 CFR | Part 11, Part 210 |
| GAMP5 | Guide for Validation of Automated Systems |
| URS | URS for Biometric MES System |
| FS | NYMI FS-001 |
| DS/CS | NYMI CS-001 |

- Responsibilities

  - System Owner

To provide all required documentation, create the performance qualification testing document, and to provide a functional test environment

  - System Validation Tester

To perform the qualification testing and record the results in this document and documenting any issues that are encountered

- System Description

Unless otherwise noted, NES and AD servers reside in the same domain as do the user computers

Each User computer has one attached Nymi-supported NFC reader and one attached Bluegiga BLE adapter

The configuration falls into GAMP5 Category 3

- Test Procedure

| System Name: | Nes-and-evidandomain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing NES operations |  |  |  |
| Purpose | Validate that the NES is functional in the IT Infrastructure and that a user can enroll in Nymi Band |  |  |  |
| Test Reference: |  | OQ-001 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-001, URS-002 Functional Specification: FS-CFG-001, FS-CFG-010 |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Connect to the NES Administration console | A secure NES Console webpage appears | NES Console Login page appears | Pass |
| 2 | Log into the NES Administration Console with an administrator account | Log in succeeds | Login with administrator username and password succeeds | Pass |
| 3 | Navigate to the About page and click View Full System Diagnostics | The System Diagnostics page does not report errors | All items in the System Diagnostics page report "Pass" | Pass |

| System Name: | Nes-and-evidiandomain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing operations in a multi-domain environment |  |  |  |
| Purpose | Validate that the Nymi Evidian solution is functional in a multidomain environment, users in different domains can enroll Nymi Band to the same Evidian instance, and NES administrators in different domains can manage NES |  |  |  |
| Test Reference: |  | OQ-002 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-003, URS-024, URS-39 Functional Specification: FS-CFG-03, FS-APP-002, FS-CFG-04 Configuration details: NES is installed in domain A Users are in domain A, B, C Domain A is a root domain and B is A's sub-domain Domain C is the root domain of another forest There is a two-way trust from A to C |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Use a computer and user account on Domain A to enroll a Nymi Band using Nymi Band Application | Enrollment succeeds | Enrollment succeeds | Pass |
| 2 | Use a computer and user account on Domain B to enroll a Nymi Band to using Nymi Band Application | Enrollment succeeds | Enrollment succeeds | Pass |
| 3 | Use a computer and user account on Domain C to enroll a Nymi Band using Nymi Band Application | Enrollment succeeds | Enrollment succeeds | Pass |
| 4 | Connect to the NES Administration console from a terminal on Domain A, B or C | A secure NES Console webpage appears | NES Console login page appears | Pass |
| 5 | Log into the NES Administration Console with an administrator account in Domain A | Login succeds | Log in succeeds | Pass |
| 6 | Open Policies page to view policies | Policies pages opens successfully | Policy page appears | Pass |
| 7 | Open the default policy and edit and save the changes | Policy saves succeeds | Policy successfully saved | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing that the solution supports multiple authentication methods |  |  |  |
| Purpose | Validate that Nymi Enterprise Edition supports authentication by AD credentials and/or fingerprint verification |  |  |  |
| Test Reference: |  | OQ-003 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-030 Functional Specification: FS-NB-015, FS-APP-001 |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Instruct a user to use Nymi Band Application to enroll a Nymi Band | Enrollment succeeds | Nymi Band displays a checkmark (successful authentication) and battery sign | Pass |
| 2 | Instruct the user to remove Nymi Band | Nymi Band de-authenticates | Nymi Band vibrates and displays the user label(name) and battery sign | Pass |
| 3 | Instruct the user to place Nymi Band back on the same wrist that was used during the enrollment process | Nymi Band prompts the user to authenticate | Nymi Band displays a fingerprint sign prompting the user to authenticate | Pass |
| 4 | Instruct the user to place the finger that they used to enroll Nymi Band on the fingerprint sensor for the amount of time indicated on Nymi Band screen | Nymi Band authentication succeeds | Nymi Band displays a checkmark and battery sign | Pass |
| 5 | Instruct the user to remove Nymi Band | Nymi Band deauthenticates | Nymi Band vibrates and displays user label and battery sign on | Pass |
| 6 | Instruct the user to log in to Nymi Band Application using their corporate credentials | Login succeeds | Log in succeeds | Pass |
| 7 | Instruct the user to click the Authenticate button | Nymi Band authentication succeeds | Nymi Band is authenticated, and the band displays a checkmark and battery sign | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing solution interactions with Active Directory |  |  |  |
| Purpose | Validate that Nymi Connected Worker Platform solution ensures that Nymi Band user is valid in Active Directory Usernames and passwords are not stored by NES |  |  |  |
| Test Reference: |  | OQ-004 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-013 Functional Specification: FS-NB-016 |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Instruct a user to use Nymi Band Application to enroll a Nymi Band | Enrollment succeeds | Nymi Band displays a checkmark (successful enrollment) and battery sign | Pass |
| 2 | Instruct the IT admin to log in to a DC and disable the AD account for the user that is associated with Nymi Band | Disabling the user account succeeds | The account is disabled | Pass |
| 3 | Instruct the user to remove Nymi Band and put it back on | Nymi Band deauthenticates | Nymi Band displays the user label and battery sign | Pass |
| 4 | Instruct the IT admin to log into a workstation | Login is successful | Log in succeeds | Pass |
| 5 | Instruct the user to log into Nymi Band application on the workstation | Log in fails with an error | Log in fails | Pass |
| 6 | Instruct the IT admin to re-enable the user’s AD account | Enabling the user account succeeds | Account is enabled | Pass |
| 7 | Instruct the user to log into Nymi Band Application | Login is successful in Nymi Band Application | Nymi Band Application displays the Authenticate button | Pass |
| 8 | Instruct the user to click the Authenticate button | Authentication succeeds | Nymi Band is authenticated and displays a checkmark and battery sign | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing that the Nymi solution has an interface for enrollment |  |  |  |
| Purpose | Validate that Nymi Band application provides users with an interface to associate Nymi Bands with users |  |  |  |
| Test Reference: |  | OQ-005 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| URS tested: URS-026, URS-030, URS-039 Functional specification: FS-NB-015, FS-APP-001, FS-PHY-007 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Locate Nymi Band Application and follow the steps on the interface to enroll your Nymi Band | Enrollment succeeds | Nymi Band displays a checkmark (successful enrollment) and battery sign | Pass |
| 2 | Take out Nymi Band from the user’s wrist | Nymi Band de-authenticates | Nymi Band vibrates and displays the user label(name) and battery sign | Pass |
| 3 | Put Nymi Band on your wrist | Nymi Band prompts the user to authenticate | Nymi Band displays a fingerprint icon prompting the user to authenticate | Pass |
| 4 | Log in to Nymi Band Application using the same credentials that you used when you enrolled your Nymi Band | Log in succeeds | Nymi Band Application displays the Authenticate button | Pass |
| 5 | Authenticate Nymi Band by pressing the “Authenticate” Button in Nymi Band Application | Authentication succeeds | Nymi Band authenticates and | Pass |
| 6 | Take out Nymi Band from the user’s wrist | Nymi Band de-authenticates | Nymi Band vibrates and displays the user label with battery sign | Pass |
| 7 | Put Nymi Band on your wrist | Nymi Band prompts the user to authenticate | Nymi Band displays a fingerprint icon prompting the user to authenticate | Pass |
| 8 | Put the finger that you used to enroll Nymi Band on the fingerprint reader and follow the instructions to authenticate Nymi Band | Authentication succeeds | Nymi Band is authenticated with a successful fingerprint | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Test removal of biometric information from Nymi Band |  |  |  |
| Purpose | Validate that Nymi Connected Worker Platform provides users with the ability to remove biometric information Nymi Band |  |  |  |
| Test Reference: |  | OQ-006 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: USR-012 Functional specification: FS-NB-012 |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Instruct a user to use Nymi Band Application to enroll a Nymi Band | Enrollment succeeds | Nymi Band displays a checkmark and battery sign | Pass |
| 2 | Instruct the user to remove Nymi Band | Nymi Band de-authenticates | Nymi Band vibrates and the band displays user label and battery sign | Pass |
| 3 | Instruct the user to place Nymi Band back on the same wrist that was used during the enrollment process | Nymi Band prompts the user to authenticate | Nymi Band displays a fingerprint icon prompting the user to authenticate | Pass |
| 4 | Instruct the user to place the finger that they used to enroll Nymi Band on the fingerprint sensor for the amount of time indicated on Nymi Band screen | Nymi Band authentication succeeds | Nymi Band is authenticated with a successful fingerprint | Pass |
| 4 | Instruct the user to remove Nymi Band | Nymi Band deauthenticates | Nymi Band vibrates and displays the user label and battery sign | Pass |
| 5 | Instruct the user to plug Nymi Band in to charge | Nymi Band starts to charge | Nymi Band screen displays the charging icon | Pass |
| 7 | Instruct the user to hold the bottom button of Nymi Band until the Deleting User Data appears | Nymi Band data deletion succeeds | Nymi Band screen displays “Deleting User Data”, boot sequence information and then displays No User and battery sign | Pass |
| 8 | Instruct the user to put on Nymi Band | Nymi Band does not prompt the user to authenticate | Nymi Band displays the setup code | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing firmware updates |  |  |  |
| Purpose | Validate that firmware updates can be performed concurrently on up to 5 Nymi Bands and without Nymi assistance |  |  |  |
| Test Reference: |  | OQ-007 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-025, URS-009 Functional specification: FS-BAT-006, FS-BAT-005 |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Obtain a USB charging hub that has at least 10 ports | n/a | n/a | n/a |
| 2 | Plug 5 Nymi Bands and 5 Bluetooth adapters into the USB hub and ensure that the hub is plugged into a power supply | Nymi Bands start to charge | Nymi Bands display the charging symbol | Pass |
| 3 | Place the hub within Bluetooth range of a computer that has Nymi Band Application software installed and a Bluetooth adapter plugged into it | n/a | n/a | n/a |
| 4 | From a command prompt on the computer, type the command to start the firmware updater script | The firmware updater script starts, and Nymi Band displays a message to indicate that the update is in progress | The firmware updater displays a status window for the firmware updates and Nymi Bands screens display “STANDBY” | Pass |
| 5 | While Nymi Band is displays "STANDBY", take Nymi Band off charge | Firmware update stops and Nymi Band does not display an error | Nymi Band does not display "STANDBY" or report an error | Pass |
| 6 | Leave Nymi Band off charge for 5 mins | The firmware update does not start | Nymi Band does not display "STANDBY" | Pass |
| 7 | Put Nymi Band back on charge | The update of Nymi Band firmware starts again | Nymi Band displays "STANDBY" | Pass |
| 8 | Monitor Nymi Band display and the firmware updater status window | Nymi Band firmware updates complete and the administrator is provided with visual indications | Nymi Bands reboot as each update completes and the status window provides information about the number of completed Nymi Bands | Pass |
| 9 | In Windows Explorer, go to the Properties of the firmware updater script file | The properties of the file provide signing information | Signing information appears on the Digital Signatures tab | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing Nymi Band cleaning |  |  |  |
| Purpose | Validate that Nymi Band can be sanitized with an alcohol wipe or spray |  |  |  |
| Test Reference: |  | OQ-008 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-008 Functional specification: FS-ENV-003 |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Using a wipe that has been soaked in 70% isopropanol, wipe all surfaces of Nymi Band’s body including: -Front surface (display, fingerprint sensor) -Back surface (ECG electrode, sensor glass) -Sides (button surfaces, USB port) -exposed strap surfaces | Nymi Band is cleaned and the body of Nymi Band is undamaged by cleaning | Nymi Band displays battery symbol and the body appears unchanged | Pass |
| 2 | Wear Nymi Band on your wrist and authenticate it using your fingerprint | Nymi Band prompts the user to authenticate, and then authentication succeeds | Nymi Band is authenticated following successful fingerprint | Pass |
| 3 | Remove Nymi band off the wrist | Nymi Band deauthenticates | Nymi Band vibrates,deauthenticates and then displays the user label and battery sign | Pass |
| 4 | Repeat these 3 steps several times throughout the day | Nymi Band deauthenticates each time | Nymi Band vibrates, deauthenticates and then displays the user label and battery sign each time | Pass |
| 5 | Place Nymi Band to charge | Nymi Band starts to charge | Nymi Band displays a battery icon that is full or has a lightning bolt | Pass |

| System Name: | Nes-and-evidian.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing Nymi-Evidian solution auditing |  |  |  |
| Purpose | Validate that the system maintains an audit log of Nymi Band user assignments in Evidian’s Enterprise Access Management (EAM) solution |  |  |  |
| Test Reference: |  | OQ-009 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-027, URS-028 Functional Specifications: FS-SAF-005 |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | In the EAM directory, search for the following objects names with objectClass “enatelUserManagedAuth”: CN=RegReqTokenId:RFIDPCSC:%NFCserialnumber% corresponding to NFC part of the band CN=RegReqTokenId:RFIDPCSC:%MACAddressBLE% corresponding to BLE part These objects are stored: Under user container in Active Directory as EAM Directory Under O=EAM/CN=Program Data/CN=Evidian Enterprise Access Management/CN=IAMForeignObjects/CN=%ActiveDirectoryUserObjectGUID% | Objects are found for user of interest | Entries found | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Test the dissociation of a user from their Nymi Band |  |  |  |
| Purpose | Validate that IT admin can dissociate a user from a Nymi Band |  |  |  |
| Test Reference: |  | OQ-010 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-019 Functional specification: FS-APP-002 |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | In the EAM Console, select the Directory panel | Directory Panel appears | Panel appears | Pass |
| 2 | Select the search request by changing the object type to user and then in the Filter field, type the username for the individual you want to disassociate Click Search | User details appear | User details window appears | Pass |
| 3 | Select the user, and then select the RFID tab | RFID Tab appears | RFID tab appears | Pass |
| 4 | Select the wearable entry and click Blacklist | Confirmation window appears and blacklisting succeeds | Entry is blacklisted | Pass |
| 5 | Select the wearable entry, and then click Delete | Confirmation window appears and wearable entry disappears | Entry is deleted | Pass |
| 6 | Select the RFID entry, and then click Delete | Delete operation succeeds | Entry is deleted | Pass |
| 7 | Select the RFID panel From the RFID state list, select Blacklisted, and then click Apply | Operation succeeds | Blacklist is successful | Pass |
| 8 | Select the RFID entry, and then click Delete | Delete operation succeeds | Entry is deleted | Pass |
| 9 | Select the Wearable entry, and then click Delete | Delete operation succeeds | Entry is deleted | Pass |
| 10 | IT admin clicks Disconnect beside Nymi Band entry in Nymi Band table | The Disconnect Nymi Band page appears | The Disconnect Nymi Band page appears | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Test the installation of an NEA on a Windows 10 thin client |  |  |  |
| Purpose | Validate that an NEA can be installed on a Windows 10 thin client |  |  |  |
| Test Reference: |  | OQ-011 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-021 Functional specification: FS-RDP-005 |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Follow the instructions in the Nymi Deployment Guide to install the Nymi Agent application on the NES host | Nymi Agent application installed successfully | Installation succeeds | Pass |
| 2 | Follow the instructions in the Nymi Deployment Guide to install the Nymi Bluetooth Endpoint application on the Citrix or RDP client | Installation completes successfully | Installation succeeds | Pass |
| 3 | Follow the instructions in the Nymi Deployment Guide to create the nbe.toml file on the Citrix or RDP client | File creation succeeds | nbe.toml file is in the directory | Pass |
| 4 | Follow the instructions in the Nymi Deployment Guide to install the NEA | Installation completes successfully | Installation succeeds | Pass |
| 5 | Put a Nymi Band on your wrist and authenticated it | Nymi Band prompts the user to authenticate | Authentication succeeds | Pass |
| 6 | Start the NEA | NEA prompts you for your username and password | The Login screen appears | Pass |
| 7 | Tap Nymi Band against the NFC reader | The NEA login completes successfully | Login succeeds | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing the Nymi solution with thin clients with NEAs and MES applications |  |  |  |
| Purpose | Validate that an NEA can be installed on a Windows 10 thin client and Nymi Band can be used to perform authentication tasks |  |  |  |
| Test Reference: |  | OQ-012 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-004, URS-006, URS-014, URS-015, URS-016, URS-017, URS-018, URS-020, URS-022, URS-023 Functional specification: FS-RDP-005, FS-MES-006, FS-MES-008, FS-NB-019, FS-MES-001 Configuration information: Tester is wearing a class A/C, C or D PPE suit The NFC reader is covered with 3 cm of plexiglass |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Following the steps in the Nymi Evidian Deployment Guide to install and configure the Evidian EAM controller in the environment | Installation completes successfully | Installation succeeds | Pass |
| 2 | Following the steps in the Nymi Evidian Deployment Guide to install the Evidian EAM client and Nymi Band Application for Evidian software on a machine in the environment | Installation completes successfully | Installation succeeds | Pass |
| 3 | Install the MES application Refer to the Nymi Evidian Deployment Guide for special instructions | Installation completes successfully | Installation succeeds | Pass |
| 4 | Follow the instructions in the Nymi Evidian Deployment Guide to enroll a Nymi Band | Enrollment completes successfully | Enrollment succeeds | Pass |
| 5 | Remove the Nymi Band from your wrist | Evidian cannot detect the presence of the authenticated Nymi Band | The computer locks | Pass |
| 6 | Put the Nymi Band back on your wrist and authenticate it | Nymi Band prompts the user to authenticate | Authentication completes and Nymi Band displays a checkmark (successful authentication) and battery sign | Pass |
| 7 | Tap the authenticated Nymi Band against a connected NFC reader | Log in to the computer | Log in succeeds | Pass |
| 8 | Open the MES application | The application prompts you to login | Login screen appears | Pass |
| 9 | Tap the authenticated Nymi Band against a connected NFC reader | Log in to the MES application completion | Log in succeeds | Pass |
| 10 | Perform an authentication task in the MES application | The MES application prompts you for your username and password | Username and Password prompts appear | Pass |
| 11 | Tap the authenticated Nymi Band against a connected NFC reader | The authentication task in the MES application completes successfully | MES records the authentication event | Pass |
| 12 | Walk away from the computer, until the user is outside of BLE range | Evidian cannot detect the presence of the authenticated Nymi Band | The computer locks | Pass |
| 13 | Remove the BLE adapter from the computer | Adapter removed | Adapter is on the desk | Pass |
| 14 | Following the steps in the Nymi Evidian Deployment Guide to enable NFC only | Configuration completes successfully | Configuration succeeds | Pass |
| 15 | Remove Nymi Band from your wrist | Nymi Band deauthenticates | Nymi Band displays "WEAR BAND" | Pass |
| 16 | Put Nymi Band back on your wrist and authenticate it | Nymi Band prompts the user to authenticate, and Authentication succeeds | Authentication completes and Nymi Band displays a checkmark (successful authentication) and battery sign | Pass |
| 17 | Tap the authenticated Nymi Band against a connected NFC reader. | Log in to the computer succeeds | Computer desktop appears | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing the secure Nymi Band association with user during enrollment |  |  |  |
| Purpose | Validate that the Nymi Band Application (NBA) requires the user to provide consent when enrolling a Nymi Band. |  |  |  |
| Test Reference: |  | OQ-013 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-042 Functional specification: FS-ENR-009 Configuration information: User is wearing a charged, un-enrolled Nymi Band. |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Instruct the user to open the Nymi Band Application | Nymi Band applications opens. | Nymi Band applications opens. | Pass |
| 2 | Instruct the user to tap the Nymi Band to the dongle | Prompt the user to login to the Nymi band application. | Username and Password prompts appear. | Pass |
| 3 | Instruct the user to login with their Active Directory username and password | The user entered the username and password to login. | The username and pasaword was entered and login is success. | Pass |
| 4 | Instruct the user tap the Nymi Band to the dongle to accept request consent. | The user accepts the request consent that user is assigned to this band. | The user accepts the consent for band and user association. | Pass |
| 5 | Instruct the user to perform the fingerprint capture process and authentication training. | The fingerprint print capture and authentication training are successful for this user. | The fingerprint print capture and authentication training are successful for this user. | Pass |
| 6 | The Nymi Band pushed the certs to the band and band is updated | Enrollment is completed successfully. | Enrollment is completed and the user disconnects. | Pass |

| System Name: | Nes-and-evidan.domain.com |  |  |  |
| --- | --- | --- | --- | --- |
| Test Title | Testing how to disable the haptic signals on the Nymi Band |  |  |  |
| Purpose | Validate the disabling of the haptic signals on a Nymi band. |  |  |  |
| Test Reference: |  | OQ-014 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 04-17-2025 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-040 Functional specification: FS-ENV-005 Configuration information: User is wearing a charged, unenrolled Nymi Band |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | Follow the Administration Guide and disable the haptic feedback on the Nymi band. | The haptic feedback is disabled in the policy. | The haptic feedback is disabled in the policy. | Pass |
| 2 | Follow the Administration Guide and enroll a Nymi Band. | Nymi Band is successfully enrolled without haptic feedback. | Nymi Band is successfully enrolled and does not vibrate upon completion of the enrollment. | Pass |
| 3 | Remove the Nymi band from user’s arm. | The Nymi band is removed from the user’s arm without haptic feedback. | The Nymi band is removed from the user’s arm and does not vibrate. | Pass |
| 4 | Place the Nymi Band to the user’s arm and authenticate | Nymi Band authenticated successfully without haptic feedback. | Nymi Band authenticated successfully without haptic feedback and does not vibrate. | Pass |

| Test Title | Testing Self- Service re-enrollment |  |  |  |
| --- | --- | --- | --- | --- |
| Purpose | Verify that users can re-enroll their Nymi Band while maintaining organization level access control |  |  |  |
| Test Reference: |  | OQ-016 | Test Run #: | 01 |
| Tester name: |  | Deb Claudio | Execution Date: | 06-03-2024 |

| Test Protocol |  |  |  |  |
| --- | --- | --- | --- | --- |
| URS tested: URS-039, URS-043 Functional specification: FS-APP-004, CS-APP-005 Configuration information: Self – Serve enrollment is configured and usable |  |  |  |  |
| Step# | Procedure | Expected Result | Actual Result | Pass/Fail |
| 1 | As the Site – Admin or Supervisor, follow the instructions in the Nymi Connected Worker Platform Administration Guide activate the policy to enable self-service re-enrollment policy. | Self- Service re-enrollment is enabled in the policy. | Self- Service re-enrollment is enabled in the policy. | Pass |
| 2 | As the End User, follow the instructions in the Nymi Connected Worker Platform Administration Guide to enroll a Nymi Band. | Enrollment is completed successfully. | Enrollment succeeds. | Pass |
| 3 | Open the MES application. | The application prompts you to login. | Login screen appears. | Pass |
| 4 | Tap the authenticated Nymi Band against a connected NFC reader. | The application log in completes successfully. | Log in succeeds. | Pass |
| 5 | Perform an authentication task in the MES application. | The MES application prompts you for your username and password. | Username and Password prompts appear. | Pass |
| 6 | Tap the authenticated Nymi Band against a connected NFC reader. | The authentication task in the MES application completes successfully. | MES records the authentication event. | Pass |
| 7 | As the End User, follow the instructions in the Nymi Connected Worker Platform Administration Guide to delete the user on the Nymi Band. | User deletion is completed successfully on the Nymi Band. | User deletion is completed successfully on the Nymi Band. | Pass |
| 8 | As the End User, follow the instructions in the Nymi Connected Worker Platform Administration Guide to enroll a Nymi Band. | Enrollment is completed successfully. | Enrollment is completed successfully. | Pass |
| 9 | Open the MES application. | The application prompts you to login. | Login screen appears. | Pass |
| 10 | Tap the authenticated Nymi Band against a connected NFC reader. | The application log in completes successfully. | Log in succeeds. | Pass |
| 11 | Perform an authentication task in the MES application. | The MES application prompts you for your username and password. | Username and Password prompts appear. | Pass |
| 12 | Tap the authenticated Nymi Band against a connected NFC reader. | The authentication task in the MES application completes successfully. | MES records the authentication event. | Pass |
| 13 | As the End User, follow the instructions in the Nymi Connected Worker Platform Administration Guide to delete the user on the Nymi Band. | User deletion is completed successfully on the Nymi Band. | User deletion is completed successfully on the Nymi Band. | Pass |
| 14 | As the End User, follow the instructions in the Nymi Connected Worker Platform Administration Guide to enroll a Nymi Band. | Enrollment is completed successfully. | Enrollment is completed successfully. | Pass |
| 15 | Open the MES application. | The application prompts you to login. | Login screen appears. | Pass |
| 16 | Tap the authenticated Nymi Band against a connected NFC reader. | The application log in completes successfully. | Log in succeeds. | Pass |
| 17 | Perform an authentication task in the MES application. | The MES application prompts you for your username and password. | Username and Password prompts appear. | Pass |
| 18 | Tap the authenticated Nymi Band against a connected NFC reader. | The authentication task in the MES application completes successfully. | MES records the authentication event. | Pass |